Jürgen Schulte

User’s Manual for

Bruker AM / AC

NMR Spectrometers

Available online at:

http://www.chem.binghamton.edu/staff/schulte/Manual1.htm

http://dmoz.org/Science/Chemistry/Nuclear_Magnetic_Resonance

(Netscape Open Directory Project)

Contents: 

A.		BASIC OPERATIONS	4

A.I		Locking and Shimming the Sample	4

A.IIa	Routine 1H and 13C Experiments	5

A.IIb	Steps for Manual Processing of NMR Spectra	6

A.III	Automatic Locking and Shimming	8

A.IV	Automated Procedure for Routine 1H and 13C Experiments	9

A.V		Switching between NORMAL and REVERSE mode (AM 360 only)	10

A.VI	Leaving the Spectrometer	11

B.		SPECIAL 1D EXPERIMENTS	12

B.Ia	Gated Decoupling (proton coupled carbon spectrum)	12

B.Ib		Inverse Gated Decoupling (Carbon Spectrum for Integration)	12

B.II		Homonuclear Decoupling Experiments	13

B.III	Inversion Recovery (T1 measurement)	14

B.IV	DEPT - Experiments  /   PENDANT	16

B.V		Combined DEPT Experiments	17

B.VI	NOE Difference Spectra	18

B.VII	Kinetic 1H and 13C Measurements	20

C.		2D NMR Experiments	22

C.I		Preparing the 1D projections	23

C.II		Preparing the 2D Experiment	23

C.IIa	H,H COSY	24

C.IIb	H,H NOESY	25

C.IIc	H,H TOCSY  (possible only in inverse mode)	26

C.IId	J-resolved H,H Correlation	27

C.IIe	C,H Correlation	28

C.IIf	COLOC - Long Range C,H Correlation	29

C.IIg	J-resolved C,H Correlation	30

C.IIh	INADEQUATE	31

C.III	Processing the 2D Data	32

C.IV	Plotting the 2D Spectra	33

90 Degree Pulse Lengths  - AM-360	35

AM 360  - 10 mm BB Probe Tuning	36

Emergency Shutdown of the AM-360 Spectrometer	34

Frequently Used Keystrokes	39

Table 1: Shim Matrix and Job Parameter files for Various Solvents	40

Table2: Startup Parameters for the AM-360 NMR Spectrometer	41

# A.	BASIC OPERATIONS

## A.I	Locking and Shimming the Sample

- Remove the cap from the magnet, press the orange button and then the “Lift” button.

- Remove the standard sample tube and replace it with your own cleaned sample tube. 
Press the “Lift Off” button to lower your sample into the magnet. Wait until the sample 
spins at 15-20 Hz and keep the temperature constant (300 K) during the experiment.

- Set the “LOCK POWER” and the “FIELD” to the proper values for your solvent (see reverse).

- If you are using a solvent without Deuterium, press “SWEEP OFF” and skip steps 5-9.

- Refine “FIELD” until you can cleary see the lock ringing signal.  
If you have two solvents, try both FIELD settings and use the stronger signal.

- Set the “LOCK PHASE” to 220, then press the “AUTO LOCK ” button and wait until its diode and the  “LOCK GAIN” diode have stopped flashing.

- Decrease the “LOCK GAIN” to lower the lock signal to the middle of the lock display. 
Is the lock signal too noisy? Increase the “LOCK POWER” and decrease the “LOCK GAIN”. 
Is the lock signal too strong? Decrease the “LOCK POWER” and increase the “LOCK GAIN”.

- Adjust the “LOCK PHASE” to maximize the position of the lock signal on the screen.  
If the lock signal rises off the screen, just decrease the LOCK GAIN or LOCK POWER to let it reappear and continue correcting the LOCK PHASE.

- Adjust the “Z” and “Z2” shims individually for a maximum position of the lock signal. 
Again, use LOCK GAIN or LOCK POWER to keep the signal visible. 
Repeat re-adjusting  “Z” and “Z2” several times until there is no further improvement, or:

- If your sample is concentrated enough, you can tell the spectrometer to shim for you:
Type the commands “RJ JH5D2O”, “SW5000”, “RGA”, and “TU4”. Wait.

- Press the “STDBY” button to deactivate the wheel.

- For overnight 13C experiments only:  Press “AUTO SHIM”, “Z”, “Z2”, “AUTO SHIM”.
Do not use autoshim for 1H NMR.

## A.IIa	Routine 1H and 13C Experiments

1.	Read the parameters for the experiment:

RJ  JH5CDCL3	(for a 1H experiment with a chloroform-d  solution.

PJ  			 See table 1 for other solvents.)

II

(For 13C and only on the AC-100 !!!):	Enter  MOD  1

Enter RGA 			(wait until you see “AUTO RG FINISHED…”)

Change the number of scans with NS if desired.

2.	Type:	ZG 			(wait until NS scans have been acquired)

(If the FID is clipped horizontally, decrease RG and repeat ZG)

3.	Save the FID:	WR  filename.ext  (spectrometer uses the 8+3 rule for filenames)

(To retrieve a saved file later, type: RE  filename.ext  PJ filename.ext)

4.	For 13C only: Type LB  2     EM

5.	For 1H only: Type LB  0.1    EM   SI  32   (wait for each command to finish!)

6.	Type:	FT

APK

(If the spectrum has a sinusoidal baseline, decrease RG and repeat from step 2.)

7.	Enter the Manual Processing Mode (EP ), to do manual processing (see next 2 pages).

8.	Plot the spectrum only:	PX

If PXB or PXI won’t start plotting, you will have to integrate again.

Plot the integral only:		PXI  INT1.001

Plot both (use 2 pens):	PXB INT1.001

9.	Eject the page:			NP

Notes:

–	Use two differently colored pens in the plotter.

–	Substance requirements for standard conditions:

1H  :   5 mg (MW = 500),		13C  : 50 mg (MW = 500)

## A.IIb	Steps for Manual Processing of NMR Spectra

Type “EP ” to enter the expansion and processing mode.

Phase correction (in the EP mode):

- Move the curser to the first strong signal in the spectrum. Press “P”.

- Press the vertical display (+) button, until the noise covers half a box of the grid.

- Turn knob C to properly adjust the phase of the signal at the cursor position.
(If you reach the end of the range for knob C, then press Ctrl-C to reverse it.)

- Move the spectrum to the last strong signal in the spectrum (knob A).

- Turn knob D to properly adjust the phase of this signal. Do not change C anymore! 
(If you reach the end of the range for knob D, then press Ctrl-D to reverse it.)

- Press “M” to store the phase correction into the memory.

- Hit , type ABS to correct the baseline, wait until finished, then type EP.

Calibration of the ppm scale(still in the EP mode):

- Find the solvent signal or the TMS signal, if present.

- If all the values on the screen are displayed in Hertz, press “:”.

- Place the cursor exactly on top of the signal and press “G”.

- Type the correct chemical shift value for this signal (i.e. “7.25p” for chloroform-d).

Integration of the spectrum (still in EP mode):

- Place the cursor to the left of the signal at LOWEST field.

- Press “I” to enter the integration routine. And press “Z” to mark the beginning of the integral.

- Move the cursor to the right of the signal and press “Z” to mark the end of the integral

- Repeat steps 2 to 4 for each signal of interest.

- You can normalize the integrals if you move the cursor onto the last data point of an integral, type “A” and enter a number for the area of this integral.

- After defining the last integral press “E”.

- Press ““ to accept the displayed filename or change it. This is the file which you will use to print out the integrals with PXB or PXI later. Remember the name!

Defining the plotting parameters (still in EP mode):

- To define the plotting range press “F” and enter low and high field limits in ppm.

- To define the minimum intensity for ‘peak picking’ place the cursor on top of the smallest signal to be picked.
Press “M” and re-enter the displayed value. Terminate with:
		“P” to pick positive peaks only (for normal 1H or 13C)
	or	“N” to pick negative peaks only
	or	““ to pick positive AND negative peaks (i.e. for DEPT)

- Either: Define the height for one particular peak: Place the cursor on top of it and press “CY” without return. Enter the height of the signal and then the length of the plot in centimeters.

- Or: Define the height for the tallest peak in the region visible on the screen: 
Press “Y” and enter the corresponding values.

- The rest of the parameters have to be entered outside of EP. Press ““ to exit.

- Set MAXX = 25, X0 = 0, Y0 = 0, CX = 25.
Set MAXY = 14 for 1H and  = 17 for 13C.
If you have NOT defined the height of the signals inside EP (Step 3 or 4), set CY = 12.

- Enter “DPO “ to define the ‘digital plotter options’. Answer all the questions so that you will receive the desired output of your spectra. 
Notes: -if you want to print the integration, the ‘offset’ has to be  2.5 cm.
           -with CX 20 the parameters have to be printed in the upper left corner.

- Enter “PEN “ to define the colors for spectra, ppm-scale, integral, etc.
Use numbers 1 - 7 according to the position of the pens in the plotter’s caroussel.

## Printing:

1.	Plot the spectrum only:	PX

If PXB or PXI won’t start plotting, you will have to integrate again.

Plot the integral only:		PXI  INT1.001

Plot both (use 2 pens):	PXB INT1.001

2.	Eject the page:			NP

3.	To print expansions or make other changes, type EP  and start from the top of this page.

## A.III	Automatic Locking and Shimming

1.	Type: IPRS

Enter NUSO Code : H010503

Answer all questions with “” except:

enter at question “Field:”	4310 for DMSO

4400 for Methanol

4310 for Acetone

4680 for Chloroform

4490 for D2O

4680 for Benzene

4800 for Pyridine

2.	AU LOCK.AU

3.	EXE S05solvent

4.	AU SHIM.AU

5.	Wait until the message “SHIM IN PROGRESS” disappears from the top of the screen (takes 1-5 minutes).

## A.IV	Automated Procedure for Routine 1H and 13C Experiments

1.	Read the job parameter file for your solvent:

i.e.	RJ  JH5CDCL3

PJ

II

2.	Type:  SINO  4  Y

3.	Type:  AU ROUTINE.AU

The program requests a filename to save the FID.  DO NOT enter any extension!

The program will stop, when a sufficient signal-to-noise ratio has been achieved,

but may be interrupted any time by pressing CTRL-H.

4.	Make sure that the plotter has loaded paper and that there are still some pages left 
in the tray.  Use a black pen in position 1 and a differently colored pen in position 2.

5.	Type:

EXE PLOTH.EXE 		for a full 1H spectrum with integration.

EXE PLOTC.EXE 		for a full 13C spectrum.

EXE PLOTHEXP.EXE 	for expansions of a 1H spectrum.

EXE PLOTCEXP.EXE 	for expansions of a 13C spectrum.

Each program will ask for the filename of the FID.

The file must have the extension .001, but do not enter the extension.

## A.V	Switching between NORMAL and REVERSE mode (AM 360 only)

-	NORMAL  →  REVERSE

1. Type “PO  PR  0”

2. Open the metal door at the right side of the AM-360 console.

Pull out the “NORMAL MODE / INVERSE MODE” plug (upper left corner).

Turn the plug upside down and plug it back in.

(The upper label defines the current mode.)

3. At the preamplifier housing (box on the floor next to the magnet) switch the

following cables:

a) Connect the cable labeled “Preamp.2” with the cable “F1IN” using a BNC connector.

b) Plug the BNC cable “F2IN” into the plug “Transm. F1” at the side of the box.

4. Type “PR  H1”

The NORMAL 1H and 13C experiments are NOT possible in reverse mode!

You need a special microprogram to detect 1H.

-	REVERSE  →  NORMAL

1. Type “PO  PR  0”

2. Open the metal door at the right side of the AM-360 console.

Turn the plug “NORMAL / REVERSE”, so that the label “NORMAL” is on top.

3. a) Plug the cable “F1IN” into the “TRANSM. F1” port on the preamplifier housing box.

b) Plug the cable “F2IN” into the “DECOUPLER IN” port.

c) Plug the cable “Preamp.2” into the “PREAMP.2-SELECTIVE” port.

4. Type “PR  11”

## A.VI	Leaving the Spectrometer

1.	Put the D2O sample into the magnet.

2.	Type:  EXE FINISH

3.	Verify the settings of the Temperature unit:

Temperature:  300 K

Heater Power Limiter:  5 to 6

4.	Please remove used tissues, used plotter paper and plotter pens from the console.

Remove your samples as soon as possible from the spectrometer. Unclaimed sample

tubes and glassware will be discarded after one month.

5.	Write your experiments or other activity at the spectrometer, their duration and the status of the spectrometer into the log book.

Please report any kind of irregularities.

6.	The door to G-14 has to be locked all the time, not just during the night!

The easiest way to do this is to lock it immideately after you enter the lab, but also

check it when you leave the lab, please.

# B.	SPECIAL 1D EXPERIMENTS

## B.Ia	Gated Decoupling (proton coupled carbon spectrum)

1. Read the normal 13C parameters:

RJ .....,     PJ .....,	II

2. Type “AS GATED.AU”

and change only the following parameters:

D1 = 2,	RD = 0,	PW = 3,	P9 = 80,	S1 = 16H

NS = at least 4 times the value for 13C or “-1”

3. Start the experiment with “AU GATED.AU”.

4. The FID may be processed like a routine 13C FID.

In order to enhance the small long range couplings the FID should be multiplied with a

Gaussian-Lorentzian window-function:

set LB = -3 , GB = 0.3

enter GM

set SI = 64K

FT and process as usual.

Try using different values for GB ( between 0 and 1) to obtain the best result.

Note:	This experiment should only be used, if you want to see the C,H coupling patterns.

To find out the C,H multiplicities use the DEPT experiment (much better S/N).

## B.Ib	Inverse Gated Decoupling (Carbon Spectrum for Integration)

1. Read the normal 13C parameters:

RJ .....,     PJ .....,	II

2. Type “AS INVGATE.AU”

and change only the following parameters:

D1 = 60,	RD = 0,	PW = 3,	P9 = 80,        S1 = 16H,       NS = -1

3. Start the experiment with “AU INVGATE.AU”.

4. The FID may be processed like a routine 13C FID. Use LB = 4.

Notes:	This experiment should only be run overnight and with concentrated samples.
	If you have no quaternary carbons, you may use D1 = 30.

## B.II	Homonuclear Decoupling Experiments

1.	Perform a routine 1H experiment.

Fourier-Transform, phase correct and calibrate the spectrum properly.

2.	Type EP.

Select a signal / multiplet, which you want to decouple, place the cursor into the center

of this signal and type “O2M”. This should leave EP.

3.	Outside EP set DP = 5L and type HD.

4.	Record the FID with ZG and process the FID as usual.

Notes:

- Do not use a different power setting. This might damage the probe.

- If you want to irradiate into a signal/multiplet, which is closer than 100 Hz (0.3 ppm)

to another signal, then use DP = 8L. (only then!)

- It is better to irradiate into signals with a low multiplicity (doublet, triplet) than into a

complex multiplet.

- Result: a 1H spectrum, in which the irradiated signal has dissappeared (a tall spike may

be left) and all signals of protons which are coupling with the irradiated proton are 
   simplified due to the missing coupling.

## B.III	Inversion Recovery (T1 measurement)

a)	recording the spectra: (10 mg for 1H,  50 mg for 13C)

1. Read the routine parameters: 	RJ ....,	PJ ....,	II

2. Set PW for a 90 degree pulse (see the list at the spectrometer).

Set P2 for a 180 degree pulse, set S1 and S2 = 16H

3. Type ZG, and stop the experiment immideately with Ctrl-H. Adjust RG to avoid

clipping the height of the FID.

4. Type VD and enter ten values (customize them if necessary) into the list:

1. 0.05		2. 0.2		3. 0.4		4. 0.7		  5. 1.0

6. 2.0		7. 4.0		8. 6.0		9. 10.0		10. 20.0	11. EN

(for 13C: 0.1, 2, 5, 8, 11, 15, 20, 30, 60, 120, EN)

5. Set D1 = 30 for 1H,  120 for 13C, 		set D2 = 0.002

Set LB = 0.3 for 1H,  2 for 13C,		set RD = 0

Set NS = 8, 	NE = 10     (=the number of entries in the VD list.)

6. Type VC and enter two values into the list:

1. 10,		this must be = NE

2. 1  (for 13C: 10)

3. EN

7. Start the experiment:

AU INVREC.AU		for 1H (approx. 1 hour)

AU INVRECX.AU		for 13C (approx. 4 hours)

and enter the requested filename to store the FIDs.

b)	Processing the spectra:

1. Read the FID with the longest VD (the last one?):

RE filename.010	and	PJ filename.010

2. Set AI = 1. Confirm the re-initialization by pressing “Y”, if required.

3. Type EF, and perform a manual phase correction. Calibrate the spectrum.

4. Type  AU SPECTRA.AU

and enter two filenames:	#1: the name of the FIDs,	#2: a different name

After several minutes a set of spectra files (name: #2.001 - .010) has been created.

5. To calculate the T1 time for one particular proton, do the following:

Type VW  and switch with ‘D’ and ‘I’ to the spectrum, in which its signal has

the smallest intensity. Hit “” and “ESC” to scroll through the set of parameters,
 		until you see VD in the lower right corner. Calculate: T1 = VD / ln2.

Interpolate, if necessary.  For a more accurate method consult the next page.
c)	T1 Calculation:

1. Read the last spectrum of the series: RE  filename.010      PJ  filename.010

2. Enter EP and find the first signal of interest.

Set the cursor on top of the peak and hit “T”. Hit  ““ once to confirm the filename.

Move through the rest of the spectrum and mark the other interesting peaks with “T”.

After marking the last signal of interest, hit ““.

TURN ON THE PRINTER.

3. Type “T1”. After a few seconds the display changes. Note the function keys on the 		    bottom of the display.

4. Change “T1D” to “10”.

5. Enter “PD“ and answer the following questions:

Use Integral Point File:	N

Use EP Points File:		Y		T1PNTS

Use VD List:			N

After a minute a display of peak intenity vs. recovery delay appears for the first peak.

The data point marked with a circle can be moved with the  and  keys.

6. Type “CT1“ or hit the F2 key.

A curve will be fitted to the data points and T1 will be calculated.

7. The command “DAT1“ will calculate T1 for all remaining peaks successively.

8. If the fit for one or more of the peaks is bad, hit the F6 key until the data for this peak

are displayed on the screen. Eliminate obvious, bad fitting data points by moving the

cursor to that particular point using the  and  keys and hitting the F4 key or typing

“ELIM“.

Then recalculate T1 by hitting the F2 key.

9. To plot one particular curve, first calculate T1, then:

Set:  CX = 25,  CY = 14,  MAXY = 16

Enter “DPO“ and use the following settings:

Draw X-Axis:	Y	Offset = 1

Mark Separation: 0.5S for 1H,     5S for 13C

Peak picking:	N

Parameters: N

Rotate: N

Title: N

Start the plot with “PLTF“ or the F8 key.

10. To exit from the T1 calculation routine, type “QUIT“.

## B.IV	DEPT - Experiments  /   PENDANT

To record a single DEPT or PENDANT spectrum, please use the following procedure:

1.	Prepare sample and spectrometer as for 13C experiments.  (Lock and shim the sample.)

2.	Read the DEPT or PENDANT parameter file:

RJ   DEPTPAR.001 	(DEPT-45 gives CH3, CH2, CH:positive)

or		RJ   DEPTPAR.002 	(DEPT-90 gives CH only!)

or		RJ   DEPTPAR.003 	(DEPT-135 gives CH, CH3:positive; CH2:negative)

or		RJ   PENDANT.PAR	(gives CH, CH3: positive; C, CH2: negative)

Then type: PJ

3.	Define NS  (multiple of 32)

4.	Start the microprogram  (do not mix!!):

AU  DEPT.AU		with the DEPT parameters

AU  PENDANT.AU	for the PENDANT parameters

Notes:

- The FID has to be stored to disk by the user.

- The FID can be processed like routine 13C FIDs.

- The DEPT-135 spetrum may sometimes appear upside down (CH2 positive): 
In that case, enter NM  to flip the spectrum.

- To record all three DEPTs, please use the following procedure.  (B.V)

## B.V	Combined DEPT Experiments

a)	Starting the Experiment:

1.	Type NE  to change the duration of the experiment. (approx. 8 minutes per NE)

2.	Start the microprogram:

Type: 	AU DEPTVAR. AU 	(DEPT-45, DEPT-90 and DEPT-135)

or	AU DEPTVAR1. AU 	(3 DEPTs + 13C)

or	AU DEPTVAR2. AU 	(2 DEPTs + 13C), no DEPT-45 !!!

All programs will ask for a filename to store the FID.

The third program is recommended, because it will give an additional 13C spectrum and it won’t waste any time with the trivial DEPT-45.

b)	Processing the recorded FIDs:

1.	This procedure can only be used if one of the above experiments has been used.

It will occupy JOB 3 completely for approximately five minutes.

2.	Type: 	EXE PROCDEPT. EXE 	for the first 2 technique

EXE PROCDEP2.EXE	for the last technique

After modifying some parameters (~ 30 seconds) the spectrometer asks for the input of the FID filename.  (Enter the filename that you have used for your DEPT FIDs).  The Fourier transformed and phase corrected spectra will be stored as DEPTSPEC.001 through .004 on the hard disk.

c)	Plotting the DEPT Spectra:

1.	Make sure that there are at least five pages of plotter paper in the plotter’s tray.  One page must have been loaded and a black pen has to be in position 1 of the pen carousel.  This program will occupy JOB3 for approximately five minutes.

2.	Type: 	RE DEPTSPEC.1

PJ.1

Calibrate the PPM scale of the spectrum.

3.	Type: F1  and enter the low and high field limits of your plotting region.  (The default setting will print the whole spectrum from +220 ppm to -20 ppm.)

4.	Type:  EXE PLOTDEPT.EXE  or  EXE PLOTDEP2.EXE   (see above)

After approximately 1 minute the plot will start from JOB3 and will be finished after 5 minutes.  The three DEPT spectra will be plotted on a single page and the optional 13C spectrum on a separate page.

5.	After the plot is finished, you can turn on the printer and enter “PP  ” to print out the peak positions with the CH multiplicity label.

## B.VI	NOE Difference Spectra

a)	Starting the Experiment:

1.	Obtain a 1H spectrum with standard parameters.

2.	Enter “EP  ”

and set the cursor into the center of the largest region without any peaks (reference point). Press “O2L” and accept the displayed filename (FQLIST.001) by pressing .

3.	Place the cursor on top of a signal (center of a multiplet) which you want to irradiate and press “O2L”.

4.	Select further signals with your cursor and mark them each with “O2L”.

5.	Exit the EP mode by pressing “”.

6.	Set up the experiment:	AS NOEDIFF.AU

and enter the following parameters, when the program asks for them:

D1 = 0.5		D2 = 2			D3 = 0.1

S3 = 40L		RD = 0			PW = 4

NS = 8			DS = 2			NE = 100

VCLIST.001 must contain one entry:  ( = number of entries in the FQLIST)

cancel any second entry by typing “EN”.

7.	Start program with “AU  ”.

Enter a filename (without extension) for the FIDs.

b)	Processing the recorded FIDs:

1.	The following files have to be on the hard disk:

filename.001		:		reference FID (without NOE)

filename.002, etc.	:		FIDs with NOE

2.	Define L0 = number of FIDs.

3.	Type “EXE PROCNOE.EXE ”.

After the program adjusts some parameters you have to enter the filename of your FIDs and a second filename to store the difference spectra.

c)     Plotting the NOE difference spectra:

1.	Read each of the transformed difference spectra, define plotting parameters as necessary, recommended: CY = 0, MAXY = 13

2.	Enter “EP ,” press “CTRL-R” and adjust the height of the positive signals so that they are not clipped.  (Vertical Display button +/-)

3.	Press “X” to plot the spectrum.  
Write the scale factor (printed on the screen) on the plot.

d)     Calculating Quantitative NOEs:

NOE (%) =

h1 = Height or integral of a signal in the reference spectrum

h2 = Height or integral of the same signal in the difference spectrum

f1 = Scaling factor of the reference spectrum

f2= Scaling factor of the difference spectrum

Notes:

–	Integration of signals is more accurate than measuring the heights.

–	NOEs may be negative (very rare) or zero.

–	NOE experiments should be run overnight, but not for more than five hours.  To minimize artifacts, turn off the air conditioning and the sample spinning.

## B.VII	Kinetic 1H and 13C Measurements

1.	Define L0:

The total number of scans will be (NS cannot be changed):

L0*8 scans for each 1H experiment.

L0*1024 scans for each 13C experiment.

2.	Define L1:

This is the waiting time before the 1H and 13C experiments will be repeated.  Enter the 	value in seconds.

3.	Define L2:

Enter the number of 1H and 13C experiments to be recorded.

6.	Type AU KINETIC. AU

and enter the requested filename without extension.

Notes:

–	The program will perform a 1H and a 13C routine experiment with L0*8 and L0*1024 scans, then wait for L1 seconds and repeat these steps L2 times.

–	L0 = 1 is sufficient for a sample with more than 20 mg of your substrate.

–	The total duration of the experiment is:

[(1000*L0) + L1] * L2 seconds.

## B.VIII	Hahn-Echo (T2 measurement)

a)	recording the spectra: (10 mg for 1H,  50 mg for 13C)

1. Read the routine parameters: 	RJ ....,	PJ ....,	II

2. Set P1 for a 90 degree pulse (see the list at the spectrometer).

Set P2 for a 180 degree pulse. Set RD and PW = 0

3. Type ZG, and stop the experiment immideately with Ctrl-H. Adjust RG to avoid

clipping the height of the FID.

4. Type VD and enter ten values (customize them if necessary) into the list:

1. 0.05		2. 0.2		3. 0.4		4. 0.7		  5. 1.0

6. 2.0		7. 4.0		8. 6.0		9. 10.0		10. 20.0	11. EN

(for 13C: 0.1, 2, 5, 8, 11, 15, 20, 30, 60, 120, EN)

5. Set D1 = 30 for 1H,  120 for 13C

Set LB = 0.3 for 1H,  2 for 13C

Set NS = 8, 	NE = 10     (=the number of entries in the VD list.)

6. Start the experiment:

AU HAHNECHO.AU

and enter the requested filename to store the FIDs.

b)	Processing the spectra:

1. Read the FID with the shortest VD (the first one?):

RE filename.001	and	PJ filename.001

2. Set AI = 1. Confirm the re-initialization by pressing “Y”, if required.

3. Type EF, and perform a manual phase correction. Calibrate the spectrum.

4. Type  AU SPECTRA.AU

and enter two filenames:	#1: the name of the FIDs,	#2: a different name

After several minutes a set of spectra files (names: #2.001 - .010) has been created.

5. To calculate the T2 time for one particular proton, do the following:

Type VW  and switch with ‘D’ and ‘I’ to the spectrum, in which its signal has

vanished. Find VD for this spectrum and T2 = VD / ln2.

(A more accurate T2 calculation can be done with the T1/T2 software on the spectrometers.

See section B.IIIc and use CT2 instead of CT1 to calculate T2.

Be careful! If you type CT1 by mistake, the computer will most likely freeze up.)

# C.	2D NMR Experiments

Since 2D Experiments are more difficult to set up, than 1D Experiments, every user should be

extremely cautious when attempting to start the procedure.

The preparation of the experiments requires a lot of interaction of the user with the spectrometer.

If you are not 100% certain about the procedure, you should ask the NMR Specialist for assistance. Otherwise you might do serious damage to the equipment.

The following list of experiments is currently available and will be continuously expanded:

(The substrate requirements show the amount of sample needed, to obtain a clean spectrum

in the indicated time period. A molecular weight of 500 is assumed.)

a.	H,H COSY:			to identify protons, which couple with each other

Substrate requirement:	5 - 10 mg,	30 minutes

b.	H,H NOESY:			to identify the distance between different protons

Substrate requirement:	10 - 20 mg,	4 hours

c.	H,H TOCSY:			to see the complete coupling networks between protons

Substrate requirement:	5 - 10 mg,	1 hour

d.	J-resolved H,H correlation:	to separate overlapping 1H multiplets

Substrate requirement:	5 - 10 mg	1 hour

e.	C,H Correlation:		to identify directly connected protons and carbon atoms

Substrate requirement:	50 - 100 mg	1 hour

f.	COLOC:			long range (2,3 bonds) C,H correlation

Substrate requirement:	50 - 100 mg	8 hours

g.	J-resolved C,H Correlation:	to separate overlapping 13C multiplets

Substrate requirement:	50 - 100 mg	1 hour

h.	C,C INADEQUATE:	to identify directly connected carbon atoms

Substrate requirement:	>1000 mg	>24 hours

All 2D experiments can be broken down into 4 steps:

I.	Preparing the 1D projections (These will be used when you plot the 2D spectra.)

II.	Preparing the 2D experiment (Enter all parameters for the 2D experiment.)

III.	Processing the 2D data set.

IV.	Plotting the 2D spectrum.

All the steps have to be followed precisely and in the order in which they are listed!

## C.I	Preparing the 1D projections

1.	Record a routine Proton NMR Spectrum. Process it as usual (Fourier transformation,

phase correction, baseline correction, calibration). Set CX = 15, CY = 15, MAXY = 17.

2.	Write the spectrum to the disk. (Use only numbers in the extension, no letters!)

(You have to save it NOW, NOT later!)

3.	Enter the EP mode and press Ctrl-R to display the complete spectrum.

Set the cursor in front of the first signal in the spectrum (you may ignore a solvent signal,

if it is far away from the signals of your compound.) Press R to define the start of the

relevant spectrum. Move the cursor behind the last signal in the spectrum and press R

again to define the end. Now the display should only show the desired region.

Be careful, not to exclude any signals from your compound!

Do NOT press “enter”, or you will have to repeat step 3.

4.	At this point press Ctrl-O. This will leave the EP mode and display the changed values

for O1 and SW. Type SW and increase its value by approx. 20%.

Write down O1, SW and SR, you will need them later.

This spectrum must not be stored to disk, as it is useless!!

5.	Repeat steps 1 to 4 for Carbon-13, you should change CX  to 18 in step 1.

( The C-13 projection is only necessary for those 2D experiments, which involve the

C-13 nucleus like the C,H correlations and the INADEQUATE experiment.)

6.	Continue with the preparation of the individual 2D experiments.

## C.II	Preparing the 2D Experiment

ONLY use the procedure, which describes YOUR experiment. Please follow it exactly!

Now you will need your list with the O1, SW and SR values from step C.I.4.

### C.IIa	H,H COSY

1.	Read your PROTON projection: RE  filename.extn   PJ  filename.extn

2.	Set the following parameters:	NOBC = 0

SI = 1K,	ND0 = 1,	MC2 = M,	RG = 8,	WDW1 = S,	WDW2 = S,	SSB1 = 3,	SSB2 = 3, 	SR1 and SR2 = the SR from the proton spectrum.

O1 = O2 = O11 = the O1 you found for the proton spectrum.

SW2 = the SW you found for the proton spectrum,	SW1 = half of SW2.

3.	Set up the COSY experiment:  AS  COSYDQF.AU

This will display the complete microprogram and then (maybe after a few “  ”)

it will print out the parameters it needs on the LCD display or on the screen.

You may accept them by hitting return or you may change them one by one by entering

your new values. Recommended are:

D1 = 2		P1 = 12.5	D0 = 3U	D2 = 50M

D3 = 3U	RD = 0		PW = 0	DE must not be changed!

NS = 16	DS = 2		NE = 256	IN must not be changed!

Please keep in mind that you must not press “  “ after pressing a letter key!

If you do so by mistake, press Ctrl-Q and repeat step 3.

4. 	Type ST2D     to bring up the 2D parameter screen.

Adjust SI and NE simultaneously (only values of 2n are allowed here) and repeat

step 4 until the resolution (Hz/Pt) is between 3 and 6.

If you need a better resolution, there is a better experiment.(Ask the NMR specialist.)

5.	Type EXPT   to display the duration of the experiment.

If the duration is too long, decrease D1 (but not below “1”) or NE (decrease not by more

than 25%, even numbers only) and check the duration again until it is acceptable.

Do not type ST2D after changing NE!!

You may also change NS to 8, if the sample is concentrated enough (>10 mg) or to

multiples of  16 if the sample is very dilute (<2mg).

6.	Check I2D, if it is not exactly = 1.000 then you must start over from step 1!!

7.	Check SF1 and SF2. If they are not equal to the proton frequency, start over from step 1!!

8.	Write the parameters to the disk: WJ2D filename.2DP

9.	Turn off sample spinning and increase the “LOCK GAIN” to the top edge of the screen.

10.	Start the experiment: AU COSYDQF.AU

Enter the requested filename. The extension “.SER” is required!!

Watch the FID during the first scan. If it is too big (1 box from the center) or too small,

press “Ctrl-H”, change the setting for RG and restart the experiment with AU.

### C.IIb	H,H NOESY

1.	Read your PROTON projection: RE  filename.extn   PJ  filename.extn

2.	Set the following parameters:

SI = 1K,	ND0 = 2,	MC2 = W, 	REV = Y,	REDF = N,
	WDW1 = S,	WDW2 = S,	SSB1 = 3,	SSB2 = 3,	RG = 8,	NOBC = 1

O1 = O2 = O11 = the O1 you found for the proton spectrum.

SR1 and SR2 = the SR you found for the proton spectrum.

SW2 = the SW you found for the proton spectrum,	SW1 = half of SW2.

3.	Set up the NOESY experiment:  AS  NOESYPH.AU

This will display the complete microprogram and then (maybe after a few “  ”)

it will print out the parameters it needs on the LCD display or on the screen.

You can change them one by one by entering the new values. Recommended are:

D1 = 2		P1 = 12.5	D0 = 3U	P2 = 12.5

D9 = 2		P3 = 12.5	RD = 0		PW = 0	DE must not be changed!

NS = 16	DS = 2		NE = 512	V9 = 1		IN must not be changed!

Please keep in mind that you must not press “  “ after pressing a letter key!

If you do so by mistake, press Ctrl-Q and repeat step 3.

4. 	Type ST2D     to bring up the 2D parameter screen.

Adjust SI and NE simultaneously (only values of 2n are allowed here) and repeat

step 4 until the resolution (Hz/Pt) is between 3 and 6.

5.	Type EXPT   to display the duration of the experiment.

If the duration is too long, decrease D1 (but not below “1”) or NE (decrease not by more

than 25%, even numbers only) and check the duration again until it is acceptable.

Do not type ST2D after changing NE!!

It is not advisable to change NS, but if you do, then only in multiples of 8.

6.	Check I2D, if it is not =1.000 then you must start over from step 1!!

7.	Check SF1 and SF2. If they are not equal to the proton frequency, start over from step 1!!

8.	Write the parameters to the disk: WJ2D filename.2DP

9.	Turn off the sample spinning and increase the “LOCK GAIN” until the lock signal is at

the top edge of the screen.

10.	Start the experiment: AU NOESYPH.AU

Enter the requested filename. The extension “.SER” is required!!

Watch the FID during the first scan. If it is too big (1 box from the center) or too small,

press “Ctrl-H”, change the setting for RG and restart the experiment with AU.

### C.IIc	H,H TOCSY  (possible only in inverse mode)

0.	Switch the spectrometer to inverse operation mode. (Ask the NMR specialist for help.)

1.	Read your PROTON projection: RE  filename.extn   PJ  filename.extn

2.	Set the following parameters:

SI = 2K,	ND0 = 2,	MC2 = W, 	REV = Y,	REDF = N,
	WDW1 = S,	WDW2 = S,	SSB1 = 6,	SSB2 = 6,	RG = 8,	NOBC = 1

O1 = O2 = O11 = the O1 you found for the proton spectrum.

SR1 and SR2 = the SR you found for the proton spectrum.

SW2 = the SW you found for the proton spectrum,	SW1 = half of SW2.

3.	Set up the TOCSY experiment:  AS  TOCSY2D.AU

This will display the complete microprogram and then (maybe after a few “  ”)

it will print out the parameters it needs on the LCD display or on the screen.

You can change them one by one by entering the new values. Recommended are:

D1 = 2		S1 = 8H	P1 = 38	D0 = 3U	P3 = 2500

P2 = 76	L6 = 50	P4 = 2500	RD = 0		PW = 0

DE must not be changed!	NS = 16	DS = 2		NE = 1024

IN must not be changed!

Please keep in mind that you must not press “  “ after pressing a letter key!

If you do so by mistake, press Ctrl-Q and repeat step 3.

4. 	Type ST2D     to bring up the 2D parameter screen.

Adjust SI and NE simultaneously (only values of 2n are allowed here) and repeat

step 4 until the resolution (Hz/Pt) is between 3 and 6.

5.	Type EXPT   to display the duration of the experiment.

If the duration is too long, decrease D1 (but not below “1”) or NE (decrease not by more

than 25%, even numbers only) and check the duration again until it is acceptable.

Do not type ST2D after changing NE!!

It is not advisable to change NS, but if you do, then only in multiples of 8.

6.	Check I2D, if it is not =1.000 then you must start over from step 1!!

7.	Check SF1 and SF2. If they are not equal to the proton frequency, start over from step 1!!

8.	Write the parameters to the disk: WJ2D filename.2DP

9.	Turn off sample spinning and increase the “LOCK GAIN” to the top edge of the screen.

10.	Start the experiment: AU TOCSY2D.AU

Enter the requested filename. The extension “.SER” is required!!

Watch the FID during the first scan. If it is too big (1 box from the center) or too small,

press “Ctrl-H”, change the setting for RG and restart the experiment with AU.

### C.IId	J-resolved H,H Correlation

1.	Read your PROTON projection: RE  filename.extn   PJ  filename.extn

2.	Set the following parameters:

SI = 4K,	ND0 = 2,	MC2 = M, 	REV = N,	REDF = N,
	WDW1 = S,	WDW2 = S,	SSB1 = 4,	SSB2 = 4,	RG = 8,	NOBC = 0

O1 = O2 = O11 = SR1 =  the O1 you found for the proton spectrum.

SR2 = the SR you found for the proton spectrum.

SW2 = the SW you found for the proton spectrum,	SW1 = ‘30.0’.

3.	Set up the J-resolved experiment:  AS  JRES.AU

This will display the complete microprogram and then (maybe after a few “  ”)

it will print out the parameters it needs on the LCD display or on the screen.

You can change them one by one by entering the new values. Recommended are:

D1 = 2		P1 = 12.5	D0 = 3U	P2 = 25.0

RD = 0		PW = 0				DE must not be changed!

NS = 16	DS = 2		NE = 64		IN must not be changed!

Please keep in mind that you must not press “  “ after pressing a letter key!

If you do so by mistake, press Ctrl-Q and repeat step 3.

4. 	Type ST2D     to bring up the 2D parameter screen.

Adjust SI (only values of 2n are allowed here) and repeat step 4 until the resolution

(Hz/Pt)  in the F2 dimension is between 0.5 and 1.5.

5.	Type ST2D  and then type EXPT   to display the duration of the experiment.

If the duration is too long, decrease D1 (but not below “1”) or NE (decrease not by more

than 25%, even numbers only) and repeat step 5  until the duration is acceptable.

It is not advisable to change NS, but if you do, then only in multiples of 8.

6.	Type I2D  and enter the 2n-number closest to the displayed value.

(i.e. set it to . . ., 0.25, 0.5, 1, 2, 4, 8, 16, . . .; only 2n-values are allowed!)

7.	Double SI2 and set SI1 to ‘256W’ to improve the resolution.

8.	Check SF1 and SF2. If they are not equal to the proton frequency, start over from step 1!!

9.	Write the parameters to the disk: WJ2D filename.2DP

10.	Turn off the sample spinning and increase the “LOCK GAIN” until the lock signal is at

the top edge of the screen.

11.	Start the experiment: AU JRES.AU

Enter the requested filename. The extension “.SER” is required!!

Watch the FID during the first scan. If it is too big (1 box from the center) or too small,

press “Ctrl-H”, change the setting for RG and restart the experiment with AU.

### C.IIe	C,H Correlation

1.	Read your CARBON-13 projection: RE  filename.extn   PJ  filename.extn

2.	Set the following parameters:

SI = 1K,	ND0 = 2,	MC2 = M, 	REV = N,	REDF = N,
	WDW1 = S,	WDW2 = S,	SSB1 = 4,	SSB2 = 4,	NOBC = 0

O2 = O11 = the O1 you found for the proton spectrum.

O1 = the O1 you found for the carbon spectrum.

SR1 = the SR you found for the proton spectrum.

SR2 = the SR you found for the carbon spectrum.

SW1 = half of the SW you found for the proton spectrum.

SW2 = the SW you found for the carbon spectrum.

3.	Set up the C,H correlation experiment:  AS  XHCORRD.AU

This will display the complete microprogram and then (maybe after a few “  ”)

it will print out the parameters it needs on the LCD display or on the screen.

You can change them one by one by entering the new values. Recommended are:

D1 = 2		S1 = 0H	P1 = 10.4	D0 = 3U	D3 = 3.4M

P2 = 21	P4 = 10.4	P3 = 5.2	D4 = 1.7M	S2 = 16H (careful !!)

RD = 0		PW = 0	DE must not be changed!

NS = 32	DS = 0		P9 = 80	NE = 128	IN must not be changed!

Please keep in mind that you must not press “  “ after pressing a letter key!

If you do so by mistake, press Ctrl-Q and repeat step 3.

4. 	Type ST2D     to bring up the 2D parameter screen.

Adjust SI and NE (only values of 2n are allowed here) and repeat step 4 until the Hz/Pt

resolutions are between 3 and 6 in F1 and between 10 and 20 in F2.

5.	Type EXPT   to display the duration of the experiment.

If the duration is too long, decrease D1 (but not below “1”) or NE (decrease not by more

than 25%, even numbers only) and check the duration again until it is acceptable.

It is not advisable to change NS, but if you do, then only in multiples of 8.

6.	Check SF1 and SF2. If they are not equal to the 13C frequency, start over from step 1!!

7.	Change SF1 to the proton frequency (360.13).

8.	Write the parameters to the disk: WJ2D filename.2DP

9.	Turn off the sample spinning and increase the “LOCK GAIN” until the lock signal is at

the top edge of the screen.

10.	Start the experiment: AU XHCORRD.AU

Enter the requested filename. The extension “.SER” is required!!

### C.IIf	COLOC - Long Range C,H Correlation

1.	Read your CARBON-13 projection: RE  filename.extn   PJ  filename.extn

2.	Set the following parameters:

SI = 1K,	ND0 = 2,	MC2 = M, 	REV = N,	REDF = N,
	WDW1 = S,	WDW2 = S,	SSB1 = 4,	SSB2 = 4,	NOBC = 0

O2 = O11 = the O1 you found for the proton spectrum.

O1 = the O1 you found for the carbon spectrum.

SR1 = the SR you found for the proton spectrum.

SR2 = the SR you found for the carbon spectrum.

SW1 = half of the SW you found for the proton spectrum.

SW2 = the SW you found for the carbon spectrum.

3.	Set up the C,H correlation experiment:  AS  COLOC.AU

This will display the complete microprogram and then (maybe after a few “  ”)

it will print out the parameters it needs on the LCD display or on the screen.

You can change them one by one by entering the new values. Recommended are:

D1 = 2		S1 = 0H	P1 = 10.4	D0 = 3U	D3 = 50M

P2 = 21	P4 = 10.4	P3 = 5.2	D4 = 25M	S2 = 16H (careful !!)

RD = 0		PW = 0	DE must not be changed!

NS = 64	DS = 0		P9 = 80	NE = 128	IN must not be changed!

Please keep in mind that you must not press “  “ after pressing a letter key!

If you do so by mistake, press Ctrl-Q and repeat step 3.

4. 	Type ST2D     to bring up the 2D parameter screen.

Adjust SI and NE (only values of 2n are allowed here) and repeat step 4 until the Hz/Pt

resolutions are between 3 and 6 in F1 and between 10 and 20 in F2.

5.	Type EXPT   to display the duration of the experiment.

If the duration is too long, decrease D1 (but not below “1”) or NE (decrease not by more

than 25%, even numbers only) and check the duration again until it is acceptable.

It is not advisable to change NS, but if you do, then only in multiples of 8.

6.	Check SF1 and SF2. If they are not equal to the 13C frequency, start over from step 1!!

7.	Change SF1 to the proton frequency (360.13).

8.	Find out the values for NE and IN and set D3 = (NE*IN)

8.	Write the parameters to the disk: WJ2D filename.2DP

9.	Turn off the sample spinning and increase the “LOCK GAIN” to bring up the lock level.

10.	Start the experiment: AU COLOC.AU

Enter the requested filename. The extension “.SER” is required!!

### C.IIg	J-resolved C,H Correlation

1.	Read your CARBON-13 projection: RE  filename.extn   PJ  filename.extn

2.	Set the following parameters:

SI = 4K,	ND0 = 2,	MC2 = M, 	REV = N,	REDF = N,
	WDW1 = S,	WDW2 = S,	SSB1 = 3,	SSB2 = 3,	NOBC = 0

O2 = O11 = the O1 you found for the proton spectrum.

O1 = the O1 you found for the carbon spectrum.

SR1 = the O1 you found for the proton spectrum.

SR2 = the SR you found for the carbon spectrum.

SW1 = 300.

SW2 = the SW you found for the carbon spectrum.

3.	Set up the C,H correlation experiment:  AS  INEPT2D.AU

This will display the complete microprogram and then (maybe after a few “  ”)

it will print out the parameters it needs on the LCD display or on the screen.

You can change them one by one by entering the new values. Recommended are:

D1 = 2		S1 = 0H	P1 = 10.4	D2 = 1.7M	P2 = 21

P4 = 10.4	P3 = 5.2	D0 = 250U	S2 = 16H (careful !!)

RD = 0		PW = 0	DE must not be changed!

NS = 32	DS = 2		P9 = 80	NE = 128	IN must not be changed!

Please keep in mind that you must not press “  “ after pressing a letter key!

If you do so by mistake, press Ctrl-Q and repeat step 3.

4. 	Type ST2D     to bring up the 2D parameter screen.

Adjust SI and NE (only values of 2n are allowed here) and repeat step 4 until the Hz/Pt

resolutions are below 2 in F1 and between 5 and 10 in F2.

5.	Type EXPT   to display the duration of the experiment.

If the duration is too long, decrease D1 (but not below “1”) or NE (decrease not by more

than 25%, even numbers only) and check the duration again until it is acceptable.

It is not advisable to change NS, but if you do, then only in multiples of 8.

6.	Check SF1 and SF2. If they are not equal to the 13C frequency, start over from step 1!!

7.	Change SF1 to the proton frequency (360.13).

8.	Write the parameters to the disk: WJ2D filename.2DP

9.	Turn off the sample spinning and increase the “LOCK GAIN” until the lock signal is at

the top edge of the screen.

10.	Start the experiment: AU INEPT2D.AU

Enter the requested filename. The extension “.SER” is required!!

### C.IIh	INADEQUATE

1.	Read your CARBON-13 projection: RE  filename.extn   PJ  filename.extn

2.	Set the following parameters:

SI = 1K,	ND0 = 2,	MC2 = M, 	REV = N,	REDF = N,
	WDW1 = S,	WDW2 = S,	SSB1 = 3,	SSB2 = 3,	NOBC = 0

O1 = O11 = the O1 you found for the carbon spectrum.

O2 = the O1 you found for the proton spectrum.

SR2 = the SR you found for the carbon spectrum. 	SR1 = O1.

SW2 = the SW you found for the carbon spectrum.	SW1 = 0.25 * SW2

3.	Set up the C,H correlation experiment:  AS  INADSYM.AU

This will display the complete microprogram and then (maybe after a few “  ”)

it will print out the parameters it needs on the LCD display or on the screen.

You can change them one by one by entering the new values. Recommended are:

D1 = 5		P9 = 80	S1 = 18H	D3 = 5M	S2 = 16H

P1 = 5.2	D2 = 6.5M	P2 = 10.4	D0 = 2U	P3 = 7

RD = 0		PW = 0	DE must not be changed!

NS = 32	DS = 0

LO VCLIST.001

1 = 1, 2 or 4

2 = EN

NE = 256	IN must not be changed!

Please keep in mind that you must not press “  “ after pressing a letter key!

If you do so by mistake, press Ctrl-Q and repeat step 3.

4. 	Type ST2D     to bring up the 2D parameter screen.

Adjust SI and NE simultaneously (only values of 2n are allowed here) and repeat step 4 	until the Hz/Pt resolutions are around 20.

5.	Type EXPT   to display the duration of one experiment.

Multiply with the value in the VCLIST to get the total duration. If the duration is too

long, decrease NE (decrease not by more than 25%, even numbers only) or VC and check

the total duration again until it is acceptable.  NS must not be changed.

6.	Check SF1 and SF2. If they are not equal to the 13C frequency, start over from step 1!!

8.	Write the parameters to the disk: WJ2D filename.2DP

9.	Turn off the sample spinning and increase the “LOCK GAIN”to bring up the lock signal.

10.	Start the experiment: AU INADEQUATE.AU

Enter the requested filename. The extension “.SER” is required!!

## C.III	Processing the 2D Data

1.	RE filename.SER 		(read the experimental data)

2.	RJ filename.2DP 		(read the corresponding parameters

PJ filename.2DP 		  for acquisition, plotting

RJ2D filename.2DP 	  and 2D processing)

Do not change the order of these commands!

3.	The default processing parameters usually work fine, but sometimes it may be necessary

to change the window functions WDW1, WDW2 and their parameters SSB1, SSB2, LB1,

LB2, GB1, GB2, TM1, TM2 to get cleaner spectra.

4.	    Additional steps for TOCSY and NOESYPH only:

a) set NOBC = 1, type FT, enter EP and perform a manual phase correction to

make all signals negative. Exit the EP mode.

b) Type TY, enter the values 0PHZ as PC0 and 1PHZ as PC1. Type PZ.

5.	For all experiments type XFB . Depending on the size of the 2D matrix this step can

take from a few seconds to several minutes. Check the progress on the top of the screen.

6.	    Additional steps for TOCSY and NOESYPH only:

a) Type AP2D , press ‘Ctrl-B’ to see positive and negative levels and move the

cursor to a row with well separated signals. Press ‘R’ to display the row and ‘D’

or ‘I’ to de-  or increase the row number, until you find the strongest signals.

Write down this row number and memorize the shape of this 1D  spectrum.

b) Press ‘Esc’ and then ‘X’ to exit.

c) Type   RSC.SMX  ‘row number’   and correct the phase of the spectrum,

until it looks exactly like the row in AP2D.

d) Type TY, enter the values 0PHZ as PC0 and 1PHZ as PC1. Type PZ.

7.	    For TOCSY only: 				Type XF1P and then SYM

8.	    For NOESYPH only:			Type XF1P and then SYMA

9.	    For JRES only:				Type TILT  and then SYMJ

10.	    For COSY and INADEQUATE only:	Type SYMA.

11.	Type AP2D and adjust the intensity with the vertical display buttons. The correct level 	will display only cross- and diagonalpeaks and as little noise as possible.

12.	In case the calibration of one or both of the two dimensions is wrong, press ‘P’, then ‘1’

or ‘2’ for the F1 or F2 projection and then ‘E’, Press ‘Ctrl-R’ to see the whole spectrum.

Calibrate the chemical shift scale with a known signal from the 1D spectrum.

Press ‘L’ and define an expanded region with the wheels A - D. Pressing ‘X’ will expand.
	The button combination ‘M’ / ‘Y’ defines the plotting region.

13.	Press ‘Esc’ and then ‘X’ to exit AP2D.

## C.IV	Plotting the 2D Spectra

1.	Type DPO  and answer the following questions: (the bold answers are mandatory)

Draw X-Axis?		Y			Offset:		0.2

Mark Separation:	0.1P for proton, 1.0P for carbon-13

Mark cm:		- 0.2			Draw Y-axis?	Y

Mark Separation in F1 Dimension:	0.1P for proton, 1.0P for carbon-13

Parameters?		Y for homonuclear, N for heteronuclear experiments

Parameters in upper left corner?	N

Rotate?		N			Title?		Y

>	(enter the title here, up to 80 characters)

Center Title?		N

2.	Set MAXX = 25,  MAXY = 17,  X0 = 0,  Y0 = 0,  CY = 15

Set CX = 15, if the F1 and F1 dimensions are symmetrical

(COSY, NOESYPH, TOCSY, INADEQUATE)

or set CX = 18 for experiments with non-symmetrical chemical shift axes.

3.	Start the plotting sequence:

for COSY and TOCSY:		CP2P

F1 Projection:		(filename of the 1D proton spectrum)

F2 Projection:		(filename of the 1D proton spectrum)

Number of Levels:	(up to 7)		Number of Pens:	(up to 7)

Plot Frame?		Y			Grid wanted?		N

for C,H correlation and COLOC:	CP2P

F1 Projection:		(filename of the 1D proton spectrum)

F2 Projection:		(filename of the 1D carbon spectrum)

Number of Levels:	(up to 7)		Number of Pens:	1

Plot Frame?		Y			Grid wanted?		N

for NOESYPH:			C2PB

F1 Projection:		(filename of the 1D proton spectrum)

F2 Projection:		(filename of the 1D proton spectrum)

Which levels to plot:		0 (zero)

Pen no. for positive levels: 	1 (use a black pen in position 1)

Pen no. for negative levels: 	2 (use a different color in position 2)

Number of Levels:	(up to 7)

Plot Frame?		Y			Grid wanted?		N

for INADEQUATE and JRES:	CPLP

F2 Projection:		(filename of the 1D spectrum)

Number of Levels:	(up to 7)		Number of Pens:	up to 7

Plot Frame?		Y			Grid wanted?		N

for JRES: 	Set CX = 10, CY = 0, MAXY = 15 and customize DPO.

Enter AP2D and print selected columns:  press ‘C’  to select a column /

‘I’ and ‘D’ to step through the slices to find the tallest signals / ‘E’ to enter

the EP mode / ‘Ctrl-R’ to display the whole spectrum / ‘X’ to plot / .

# APPENDIX

## Emergency Procedures for the AC-300 Spectrometer

If the keyboard freezes or the computer is hanging up, please try first:

-	Ctrl-Q, then hit the RETURN key several times. Is the prompt moving? If not, then:

-	Ctrl-K (this will kill any running experiment), then hit the RETURN key several times.

-	Still nothing? Press the green START button on the console to reboot the computer.

------------------------------------------------------------------------------------------------------------------

To prevent damage to the spectrometers during a severe thunderstorm,

the following shutdown procedure has to be executed exactly in this order:

Shutdown:

- Terminate any NMR experiment with  Ctrl-H and save the data: WR filename.extn

- Type:	MO and turn off the printer, the plotter, the display and the temperature unit.

- Press “STOP” at the ASPECT 3000 computer front panel (green box).

- Switch off the hard disk(s) by turning their keys (left side of console, behind metal door).

- Switch off the ASPECT computer by turning its key.

- As the last step switch the ON/OFF button OFF (right front of AC 300 console) 
or pull the red “EMERGENCY” button (middle, front of AM-360 console).

Restart:

- Switch the ON/OFF button to ON or push the red “EMERGENCY” button back in.

- Behind the console: push the gray “MAIN POWER RESET” button to the left position.

- If the spectrometer is beeping continuously, lift up the keyboard, switch down the small “INIT” dip switch, push the black RESET button and switch the “INIT” switch back up.

- Turn on the ASPECT computer with its key and press the “STOP” button.

- Switch on all other devices (printer, plotter, display, VT unit)

- Turn on the hard disk drives (keys) and wait, until their green “READY” lights are on.

- Press “STOP    CLEAR    DISK” on the front panel of the ASPECT computer

- Confirm the date and time, which appear on the screen by pressing ““ on the keyboard.

Notes:

- After turning on the spectrometer, it takes 30 minutes for the lock circuits to stabilize.

- NEVER turn off the ASPECT computer before the hard disk. This might damage files.

## 90 Degree Pulse Lengths  -  AM-360

5 mm Dual Probe for 1H and 13C:

1. 13C - Transmitter:	90 degree =   5.2 sec

2. 1H - Transmitter:	90 degree = 12.5 sec

(used as P1 for COSY, NOESYand as PW for Inversion Recovery,

not for XHCORR and DEPT !)

3. 1H - Decoupler:

normal and reverse mode:

| DP | 90 degree pulse | Remarks: |
| --- | --- | --- |
| 0H | 10.4 sec | used as P1 in DEPT and XHCORR experiments |
| 4H | 22.0 sec |  |
| 8H | 34.0 sec | used as P1 in TOCSY and ROESY experiments |
| 14H | 78.0 sec | used as P9 for CPD decoupling (only if DP = 14H!!!) |

180 degree DANTE pulse:

100 pulses à 1.75 sec (P8)

D8 = 100 sec

DP = 8H

## AM 360 - 10 mm BB Probe Tuning

Nucleus Frequency.[MHz] Tuning Matching

2H 55.264 8494 823

7Li 139.903 8984 995

9Be 50.594 8374 803

11B 115.502 8945 992

13C 90.518 8865 966

14N 26.006 6947 755

15N 36.479 7654 822

17O 48.805 8326 800

23Na 95.227 8878 974

25Mg 22.028 6622 700

27Al 93.805 8876 975

29Si 71.515 8754 865

31P 145.731 8993 995

33S 27.611 7266 783

35Cl 35.274 7626 824

39K 16.799 5468 600

119Sn 134.180 8975 995

Limits: min. Frequency: 9.8 MHz

max. Frequency: 155.0 MHz

## 90 Degree Pulse Lengths  -  AC-300

5 mm Dual Probe for 1H and 13C:

1. 13C - Transmitter:	90 degree =   5.4 sec

2. 1H - Transmitter:	90 degree = 11.5 sec

(used as P1 for COSY, NOESYand as PW for Inversion Recovery,

not for XHCORR and DEPT !)

3. 1H - Decoupler:

normal and reverse mode:

| DP | 90 degree pulse | Remarks: |
| --- | --- | --- |
| 0H | 10.8 sec | used as P1 in DEPT and XHCORR experiments |
| 4H | 22.0 sec |  |
| 16H | 60.0 sec |  |
| 20H | 75.0 sec | used as P9 for CPD decoupling (only if DP = 20H!!!) |

180 degree DANTE pulse:

100 pulses à 1.75 sec (P8)

D8 = 100 sec

DP = 8H

## AC 300 - 5 and 10 mm BB Probe Tuning

5mm VSP 10mm VSP Nucleus Frequency [MHz] Tuning Matching Tuning Matching

2H 40.045 7765.5 844.5 7833 955

7Li 116.566 8992 984 8993 993

9Be 42.155 7678 826 7776 945

11B 96.236 8914 967 8944 985

13C 75.419 7993.5 982.5 8836 974

14N 21.600 5847 605

15N 30.394 6973 742 7457 813

17O 40.664 7647 815 7756 939

23Na 79.343 8774 944 8875 976

25Mg 18.354 4833 555 5849 666

27Al 78.158 8763 944 8855 971

29Si 59.586 7926.5 853 8666 945

31P 121.421 8997 992 8996 994

33S 23.005 6236 687 6744 735

35Cl 29.390 6917 742 7424 804

39K 13.997 2144 345 4254 515

119Sn 111.798 8975 983 8983 993

Limits: min. Frequency: 11.8 MHz 9.8 MHz

max. Frequency: 125.0 MHz 125.0 MHz

## Frequently Used Keystrokes:

A		:	continue a phase correction in EP after using APK.

Ctrl-A		:	release knob A inside EP mode.

B		:	biggest peak mode (the tallest peak will be used for any action.)

Ctrl-B		:	release knobs A and B inside EP mode.

C		:	cursor mode (the cursor position will be used for any action.)

Ctrl-C		:	invert orientation of knob C in phase correction routine.

D		:	dual display (in EP). Enter the filename of the 2nd spectrum.

Ctrl-D		: 	a) invert orientation of knob D in phase correction routine.

b) turn on/off the grid.

E		:	toggle the spectrum information (in EP).

F		:	define plotting limits (in EP).

Ctrl-F		:	show the spectrum within the plotting limits (in EP).

G		:	calibrate the chemical shift scale

Ctrl-H	or ‘Bs’	:	stop the acquisition (careful! this works everywhere!)

J		:	calculate the line width of a single peak in EP. (Answer “LORE” and .001)

Ctrl-K		:	stop all processes on the spectrometer (careful! all data will be lost!)

L		:	Load previously saved integral file (integration routine)

Ctrl-L		:	turn Lock signal and spectrum on and off

M		:	a) define the minimum intensity for peak picking (in EP).

b) start baseline correction of the integral (integration routine)

Ctrl-N		:	eject plotter paper (from inside EP)

N		:	define a Hertz/cm or ppm/cm expansion for plotting

Ctrl-O		:	delete the current command line

Ctrl-PT	:	terminate plot

Ctrl-Q		:	quit the process in the current job, does not halt a running acquisition

R		:	enlarge a region of the spectrum in EP (R, move cursor, R)

Ctrl-R		:	show the complete spectrum (in EP)

U		:	update the displayed region as new plotting region. (in EP)

X		:	plot the currently displayed screen (in EP)

Ctrl-X		:	toggle between the NMR program and a background process.

Y		:	define the tallest peak to be plotted with ......cm height (in EP)

Ctrl-Y		:	show/ hide imaginary part of the spectrum.

Z		:	mark begin and end of an integral (integration routine)

## Table 1: Shim Matrix and Job Parameter files for Various Solvents:

| Solvent | SHIM Matrix Files |  | JOB Parameter Files |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  | Proton | Carbon-13 |
| acetone-d6 | S05ACET |  |  | JH5ACET | JC5ACET |
| acetonitrile-d6 | S05CD3CN |  |  | JH5CD3CN | JC5CD3CN |
| benzene-d6 | S05C6D6 |  |  | JH5C6D6 | JC5C6D6 |
| chloroform-d | S05CDCL3 |  |  | JH5CDCL3 | JC5CDCL3 |
| DMSO-d6 | S05DMSO |  |  | JH5DMSO | JC5DMSO |
| D2O | S05D2O |  |  | JH5D2O | JC5D2O |
| methanol-d4 | S05CD3OD |  |  | JH5CD3OD | JC5CD3OC |
| methylene chloride -d2 | S05CD2CL |  |  | JH5CD2CL | JC5CD2CL |
| pyridine-d5 | S05PYR |  |  | JH5PYR | JC5PYR |
| THF-d8 | S05THF |  |  | JH5THF | JC5THF |
| toluene-d8 | S05TOLU |  |  | JH5TOLU | JC5TOLU |

## Table2: Startup Parameters for the AM-360 NMR Spectrometer

|  	  Only adjust these 3 settings:        |

*A higher LOCK POWER setting is required for concentrated samples	(more than 10 % or more than 100 mg of sample.)

## Table2: Startup Parameters for the AC-300 NMR Spectrometer

|	        Only adjust these 3 settings:           |

*A higher LOCK POWER setting is required for concentrated samples	(more than 10 % or more than 100 mg of sample.)